## Slide 1
OPTICAL METHODS FOR INVESTIGATION OF PROCESS IN PIPE

## Slide 2: OPTICAL INVESTIGATIOM METHODS FOR KINEMATIC FLOW CHARACTERISTICS
Measured data
Profile velocity (average value of two components across the tube) in four cross section along the tube
Pulsation component of velocity

## Slide 3
TRAVERSING SYSTEM
Doppler
signal
processor
Sending
optics
Laser
PC
Receiving
module
Aperture
v
VELOCIMETER SYSTEM

## Slide 4: LDA SYSTEM WITH BRAGG CELL AND LASER BEAM EXPANDER
1 - laser; 2,4 – beam splitter; 3 – Bragg cell; 5 – beam expander;
6 – frontal lens; 7 – working portion; 8 – LD receiving lens; 9 – photoelectric multiplier.

## Slide 5: BLOCK-DIAGRAM OF THE LDA DATA ACQUISITION MEASUREMENT SYSTEM
1 – light detector device; 2 – working portion; 3 – optical forming system; 4 – laser; 5 – PEM power supply; 6 - Bragg cell driver; 7-laser power supply; 8- stepper; 9 – stepper power supply and control

## Slide 6: REMOTE CONTROL ACTUATOR
1 – stepper; 2 – lead screw; 3 – guide; 4 – microswitches;
5 – scale bar; 6 – handle; 7 – bracket; 8 – movable platform;
9 – steel structure.

## Slide 7
BASIC SPECIFICATIONS OF LDA OPTICAL MODULES
Tansmission system focal distance F, mm
Level | I | II | III |  | IV
Laser type | Semiconductor | He-Ne | Semiconductor | Semiconductor | 
Laser wavelength l, μm | 0.69 | 0.6328 | 0.69 | 0.69 | 
Laser power, mW | 20 | 25 | 20 | 20 | 
Bragg cell | - | + | + | - | 
Beam expanding ratio Е | 1 | 1.94 | 1,94 | 1 | 
Transmission system focal distance F, mm | 150 | 310 | 310 | 310 | 
Light detector optics focal distance, mm | 300 | 300 | 220 | 220 | 
Beam convergence angle, degrees | 10.5 | 13.8 | 13.8 | 15.6 | 
Characteristic dimension of the measurement space 2b, μm | 131.8 | 117 | 140.4 | 272,35 | 
Fringe step, μm | 3.77 | 2.63 | 2.87 | 2.54 | 
Fringe number, ea. | 35 | 44 | 49 | 107 | 
Velocity range, m/s | +0.05 to+11.0 | -8.0 to 8.0 | -1.0 to +8.5 | +0.005 to +7.0 | 

## Slide 8: LDA SYSTEM FOR RECORDING THE VELOCITY FIELDS
1 – Steel structure;
2 – measurement set;
3– working portion;
4 – control panel;
5 – service platform

## Slide 10: OPTICAL METHODS FOR DISPERSE ANALYSIS OF TWO PHASE FLOWS
Measured data
mean size of particles (particle size distribution)
volume (mass) concentration of particles

## Slide 11
OPTICAL METHODS FOR THE FRACTIONAL ANALYSYS
Extinction method
Particle counter method (PCM)

## Slide 12: EXTINCTION METHOD (EM)
Analog to digital
converter
Sending
optics
Laser
PC
Receiver
(PM tube)

## Slide 13: PRINCIPAL OF MEASUREMENT AND CALCULATION

## Slide 14
DEPENDING OF LIGHT EXTINCTION ON MASS CONCENTRATION OF PARTICLES (r=1mm, L=100mm) (estimation)

## Slide 15
Pulse amplitude
analyzer
Sending
optics
Laser
PC
Receiving
module
PARTICLE COUNTER METHOD (PCM)

## Slide 16
COMPARISON OF EXTINCTION AND PARTICLE COUNTER METHOD
 | EXTINCTION METHOD | PARTICLE COUNTER METHOD
Principle of method | Absolute method | Calibration of the probe is necessary
Droplet sizing | Average size. Preliminary information about type of distribution is necessary (two waves method). Size distribution (multi waves method) | Size distribution. Opportunity of measuring bimodal size distribution
Concentration measurement | Measurement of “frozen” value of concentration (1/m3) | Measurement of “rate” (velocity related) concentration (1 /m 2/s). Velocity value is necessary to calculate “frozen” value of concentration
Range of measurement | 0.1-1.2 mm | 0.2-10 mm
Time resolution | On-line measurement | Measurement duration is necessary to obtain the statistic
Space resolution | 100mm3 | Less then 0.01mm3

## Slide 17: DEPOSITION THICKNESS MEASUREMENT ON THE BASE OF MEASUREMENT OF DROP’S CONCENTRATION
Velocity of deposition

## Slide 18
(estimation)
m
1mm)
mm/h